## Slide 1
Presentation on
DESIGN AND ANALYSIS OF A
TOTAL OSSICULAR REPLACEMENT PROSTHESIS (TORP)
GUIDE: DR. APURBA DAS, IIEST SHIBPUR
DEPARTMENT OF MECHANICAL ENGINEERING,
BY: SANDEEP ROY,
ROLL NO. 230710102098

## Slide 2
Introduction
“Why this matters: restoring hearing & quality of life”
  Conductive hearing loss → caused by damaged ossicles
  TORP → surgical implant to restore sound transmission
  Requirements:
1. Biocompatible material
2. Strong yet lightweight
3. No resonance in 100–6000 Hz range

## Slide 3
Objectives
  Build CAD model of TORP (Fusion 360)
  Select implant-grade material (Ti-6Al-4V)
  Perform static stress analysis under acoustic load
  Perform modal frequency analysis
  Evaluate mechanical safety & reliability

## Slide 4
CAD Design
MAIN PARTS:
  Footplate – stable contact with stapes
2. Shaft – connects base to head
3. Head – contact with the tympanic membrane
Key Features: Compact | Lightweight | Smooth transitions | Manufacturable
Footplate
Head
Shaft

## Slide 5
Ti-6Al-4V Alloy
Corrosion Resistant
High strength-to-weight ratio
Biocompatible
Excellent Fatigue & Wear Resistance
High Stiffness
MRI-Compatible & Non-Magnetic
Material Selection

## Slide 6
Simulation Methodology
Static Stress Analysis
Modal Analysis
Load = 5.0 × 10⁻⁵ N
Pressure = 1.0 × 10⁻⁶ MPa
Footplate fixed
8 eigenmodes
Goal: Avoid resonance in 100–6000 Hz
CAD
Mesh
Static
Modal
Flowchart:

## Slide 7
1
2
3
4
Max displacement: 3.339 × 10⁻⁹ mm → negligible
Max stress: 3.378 × 10⁻⁵ MPa
<< yield strength
Safety factor: 15.0 → very safe
Prosthesis behaves like rigid body
Static Analysis Results
Stress distribution plot
Displacement plot
Safety factor plot

## Slide 8
Modal Analysis Results
1
2
3
4
Natural frequencies: 7.4 kHz – 188.7 kHz
All modes lie above the 100–6000 Hz hearing band → resonance avoided
Vibrations are stable.
Only bending/twisting in higher modes
Modal frequencies of TORP – all above hearing range, with bars colored by dominant axis.
Mode 1 deformation shape at 7427 Hz, dominated by Y-axis participation.
Mode 2 deformation shape at 10,384 Hz, dominated by X-axis participation.
Mode 8 deformation shape at 188,724 Hz, dominated by X-axis participation.

## Slide 9
Static Analysis
1
Near-rigid behavior (displacement in nanometres)
2
Stress far below yield strength
3
High safety factor (15.0)
Discussion
Modal Analysis
1
Natural frequencies = 7.4 kHz to 188.7 kHz
2
All above 100–6000 Hz → no resonance
3
Stable, well-defined vibration modes
Implication
1
TORP is mechanically safe & dynamically reliable
2
Provides strong base for prototyping & trials

## Slide 10
1
2
LIMITATIONS:
Limitations & Future Work
FUTURE WORK:
  No tissue interaction/damping modeled
  Fatigue, corrosion, extreme loads not analyzed
  Simplified linear elastic model
  Nonlinear & fatigue analysis
  Prototype fabrication (CNC / 3D printing)
  In-vitro middle-ear testing
  Animal/clinical validation with surgeons

## Slide 11
CONCLUSIONS
CONCLUSION 1
Designed & simulated TORP using CAD + FEA
CONCLUSION 2
CONCLUSION 3
CONCLUSION 4
Results confirm:
  Structurally safe (negligible stress/displacement)
  Dynamically reliable (no resonance in the hearing band)
Strong candidate for prototyping & clinical validation
Internship provided valuable experience in biomedical design & simulation

## Slide 12
THANK
YOU!